# MIQE Checklist — Completed (Draft) for Supplementary Materials

This document provides a completed MIQE (Minimum Information for Publication of Quantitative Real-Time PCR Experiments) checklist for the manuscript: "Single-nucleotide primer correction significantly improves qPCR and RT-qPCR performance for adenovirus detection: a case study of the modified Jothikumar kit". 

Where information was not available from the experimental records, the field is marked as 'ACTION: required'. Please fill these fields before submission.

| MIQE Item | Response / Details |
| --- | --- |
| 1. Experimental design (biological question) | Compare performance of original (STD_AdV3) and modified (MOD_AdV3) primer sets for detection of Adenovirus 3 (ATCC VR-3™); evaluate slope, R², efficiency, LoD, and specificity. |
| 2. Sample description (source, type) | Viral template: Human adenovirus type 3 (ATCC VR-3™). Purified viral DNA used as template. No clinical samples included in this study. |
| 3. Nucleic acid extraction | NucleoSpin Dx kit (Macherey-Nagel). Manufacturer's protocol followed. 740895.50 |
| 4. Nucleic acid quantification / quality | ddPCR used to quantify stock: 3.49 × 10^6 copies/µL. Performed on Biorad QX 600. |
| 5. Primers and probes (sequences and purities) | Primers listed in Table 1. Forward STD: GGACGCCTCGGAGTACCTGAG; Forward MOD: GGACGCCTCGGAGTACCTSAG; Reverse: TTGTTCAAATCTTTGGGGTGNCA. Oligos synthesized by GENOMED, deslated. |
| 6. Primer concentrations and reaction components | Reaction volume: 10 µL; Primer concentration: 500 nM each; Master Mix: PowerUp™ SYBR™ Green Master Mix (Applied Biosystems). Template: 2 µL per reaction. |
| 7. qPCR instrument and software | BIORAD CFX Opus Dx, Software: BIORAD Maestro Dx |
| 8. Thermal cycling parameters | Initial denaturation 95 °C; 59 cycles of: 95 °C 15 s → 57 °C 16 s → 72 °C 60 s (fluorescence acquisition at extension). Melt: 65–95 °C, 0.5 °C increments. |
| 9. Reverse transcription conditions (if applicable) | Not applicable (DNA template) |
| 10. Controls (NTC, positive controls, no-RT) | No-template control (NTC) included for each run; positive control is ATCC VR-3™ template. NTC: no-amplification |
| 11. Replicates (technical/biological) | Each rection was performed in 2 technical replicates. |
| 12. Standard curve details | Ten-fold serial dilutions from 10^0 to 10^-4 leading to copies/reaction: 6.98×10^6, 6.98×10^5, 6.98×10^4, 6.98×10^3, 6.98×10^2. Linear regression of Cq vs log10(copies) used to derive slope, intercept and efficiency. |
| 13. Data analysis method | Linear regression (scipy.stats.linregress / equivalent). Efficiency E computed as (10^(-1/slope) - 1). R² from regression. Mean Cq from duplicates. Software/scripts: Python notebook (available in repository). |
| 14. Thresholding / baseline settings | Automatic. |
| 15. Melt curve analysis | Melt curves recorded and derivative (-d(RFU)/dT) analyzed to confirm single product and absence of primer-dimers. |
| 16. Specificity (in-silico and experimental) | In-silico BLAST shows modified primer aligns to multiple HAdV types and no significant matches to human genome, HSV, CMV, enteroviruses. Experimental specificity: single melt peak observed for positive reactions. |
| 17. Efficiency and linearity | MOD_AdV3: slope = -3.526, R² = 0.9998, Efficiency = 92.1% (95% CI slope: -3.61 to -3.44). STD_AdV3: slope = -5.358, R² = 0.960, Efficiency = 53.7% (95% CI slope: -8.68 to -2.04). |
| 18. Limit of Detection (LoD) and method | Preliminary LoD: MOD ≈ 7×10^2 copies/reaction (10^-4), STD ≥7×10^3 copies/reaction (10^-3). ACTION: for LoD95/LoD90 provide probit analysis input (≥20 replicates at low concentrations) if available. |
| 19. Repeatability / reproducibility | Within-run technical duplicates reported. No between-run/instrument reproducibility testing performed. ACTION: perform inter-run and inter-instrument validation for clinical application. |
| 20. Reporting of raw data | All raw Cq values, amplification and melt curve data, and scripts are available on reasonable request. |
| 21. Data transformation and statistical analysis | Cq means used for regression; slope and efficiency calculated as above. Probit analysis planned for LoD estimation if replicate data available. Statistical software: Python (scipy, statsmodels), R (optional). |
| 22. MIQE compliance statement | This checklist provides minimum required information per MIQE guidelines; where ACTION is indicated, additional metadata must be provided prior to submission to ensure full compliance. |

Prepared by: Maciej Baron, Piotr Lewandowski, Jakub Kancerek. Date: 21.11.2025